## Without Timer
TEST: TOTAL FILTRATION
Tolerance Sign | ±
Tolerance Value (kVp) | 2.0
Total Filtration Required (mm Al) | 2.5
Total Filtration At kVp | 80
mAs Station 1 | 50
mAs Station 2 | 100
mAs Station 3 | 200
Applied kVp | Measured 1 | Measured 2 | Measured 3
60 | 60.1 | 60.2 | 60.0
80 | 80.1 | 80.2 | 80.0
100 | 100.1 | 100.2 | 100.0
120 | 120.1 | 120.2 | 120.0
TEST: CONSISTENCY OF RADIATION OUTPUT
FDD (cm) | 100
Time (s) | 1.0
kVp | mAs | Measured Output 1 | Measured Output 2 | Measured Output 3 | Measured Output 4 | Measured Output 5
80 | 100 | 10.5 | 10.4 | 10.6 | 10.5 | 10.5
100 | 100 | 12.1 | 12.0 | 12.2 | 12.1 | 12.0
Tolerance Operator | <=
Tolerance Value (CoV) | 0.02
TEST: LOW CONTRAST RESOLUTION
Smallest Hole Size (mm) | Recommended Standard
1.0 | 1
TEST: HIGH CONTRAST RESOLUTION
Measured Resolution (lp/mm) | Recommended Standard (lp/mm)
2.0 | 2
TEST: EXPOSURE RATE AT TABLE TOP
Max Exposure (AEC Mode) (cGy/Min) | Max Exposure (Manual Mode) (cGy/Min) | Min. Focus to Tabletop Distance | Distance (cm) | Applied kV | Applied mA | Exposure (cGy/Min) | Mode
10 | 5 | 30 | 100 | 80 | 100 | 0.5 | AEC Mode
|  |  | 100 | 80 | 100 | 0.6 | Manual Mode
TEST: TUBE HOUSING LEAKAGE
FDD (cm) | kV | mA | Time (sec) | Workload (mA·min/week) | Tol Value | Tol Operator | Location | Front | Back | Left | Right | Top
100 | 80 | 100 | 1 | 1000 | 1.0 | <= | Tube | 0.05 | 0.03 | 0.06 | 0.04 | 0.02
[7 empty] | Collimator | 0.02 | 0.02 | 0.03 | 0.01
TEST: LINEARITY OF MAS LOADING
FDD (cm) | 100 | kV | 80
mAs | Measured Output 1 | Measured Output 2 | Measured Output 3
5 | 0.50 | 0.51 | 0.49
10 | 1.00 | 1.01 | 0.99
20 | 2.50 | 2.51 | 2.49
50 | 5.00 | 5.01 | 4.99
Tolerance Operator | <=
Tolerance Value (CoL) | 0.1