## CT Scan Data
TEST: RADIATION PROFILE WIDTH FOR CT SCAN
Applied | Measured | kVp | mA
0.6 | 0.62 | 120 | 200
1.25 | 1.27
3.0 | 3.05
TEST: MEASUREMENT OF OPERATING POTENTIAL
Set kV | Header 1 | Header 2 | Header 3 | Header 4 | Meas 1 | Meas 2 | Meas 3 | Meas 4 | Time (ms) | Slice Thickness (mm) | Tol Value | Tol Type | Tol Sign
80 | 10 | 100 | 200 | 55 | 80.1 | 80.2 | 80.3 | 78 | 100 | 5 | 5 | % | ±
100 |  |  |  |  | 100.2 | 100.1 | 100.3 | 78
120 |  |  |  |  | 120.5 | 120.4 | 120.6 | 78
TEST: MEASUREMENT OF MA LINEARITY
kVp | Slice Thickness (mm) | Time (ms) | Tolerance |  | Header 1 | Header 2 | Header 3 |  | mA Applied | Meas 1 | Meas 2 | Meas 3
80 | 5.0 | 100 | 0.1 |  | Run 1 | 78 | Run 3 |  | 10 | 10.1 | 10.2 | 10.3
[9 empty] | 20 | 20.2 | 20.1 | 20.3
[9 empty] | 50 | 50.5 | 50.4 | 50.6
[9 empty] | 100 | 101.0 | 100.8 | 101.2
TEST: TIMER ACCURACY
kVp | Slice Thickness (mm) | mA | Set Time (ms) |  | Observed Time (ms) | Tolerance (%) | Tolerance Operator
80 | 5.0 | 100 | 100 |  | 98.5 | 5 | <=
80 | 5.0 | 100 | 200 |  | 199
TEST: LINEARITY OF MAS LOADING
mAs Range | Result | FCD | kV |  | Tolerance
50 | 4.5 | 100 | 120 |  | 0.10
100 | 9.1
TEST: MEASUREMENT OF CTDI
kVp | mAs | Slice Thickness (mm) | CTDIw (Rated) Head |  | CTDIw (Rated) Body | Tol Value | Tol Sign |  | Type | Label | Head | Body
120 | 100 | 5.0 | 15.5 |  | 12.2 | 5 | ± |  | Axial Dose CTDIc |  | 16.0 | 13.0
[9 empty] | Peripheral Dose | A | 15.5 | 12.5
[9 empty] | Peripheral Dose | test | 15.7 | 12.7
[9 empty] | Peripheral Dose | C | 15.6 | 12.6
TEST: TOTAL FILTRATION
Applied KV | Applied MA | Time | Slice Thickness |  | Measured TF
120 | 100 | 1.0 | 5 |  | 7.5
TEST: Reproducibility of Radiation Output (Consistency Test)
Tolerance Sign | <=
Tolerance Value (CoV) | 0.05
mAs | Slice Thickness (mm) | Time (s) |  | Header 1 | Header 2 | Header 3 |  | Header 4 | Header 5 | kVp | Meas 1 | Meas 2 | Meas 3 | Meas 4 | Meas 5
100 | 5 | 1.0 |  | Meas 1 | Meas 2 | Meas 3 |  | Meas 4 | Meas 5 | 120 | 100.5 | 100.3 | 100.6 | 100.4 | 100.5
TEST: Radiation Leakage Level from X-Ray Tube House
kV | mA | Time (sec) | Workload |  | Workload Unit | Tol Value | Tol Operator |  | Tol Time | Location | Front | Back | Left | Right | Top
120 | 100 | 1.0 | 500 |  | mA·min/week | 1.0 | <= |  | 1 | Tube | 0.05 | 0.03 | 0.06 | 0.04 | 0.05
[10 empty] | Collimator | 0.02 | 0.02 | 0.03 | 0.01 | 0.02
TEST: LOW CONTRAST RESOLUTION
Observed Size | Contrast Level | kVp | mA |  | Slice Thickness | WW
3.0 | 1.0 | 120 | 200 |  | 5 | 400
TEST: HIGH CONTRAST RESOLUTION
Observed Size | Contrast Difference | kVp | mAs |  | Slice Thickness | WW
0.5 | 5.0 | 120 | 200 |  | 5 | 400
TEST: ALIGNMENT OF TABLE/GANTRY
Result | Tolerance
2.0 | ±2.0
TEST: TABLE POSITION
Table Position | Expected | Measured | Initial Pos |  | Load | kVp | mA |  | Slice Thickness
10 | 10 | 10.1 | 0 |  | 100 | 120 | 200 |  | 5.0
20 | 20 | 19.9
TEST: GANTRY TILT
Type | Name/Actual | Value/Measured
Parameter | Tilt Param | 45
Measurement | 0 | 0.5
Measurement | 30 | 30.2
TEST: MAXIMUM RADIATION LEVEL
Location | mR/hr | Category | Date |  | Calibrated | mA | kV |  | Time | Workload
Control Console | 0.02 | worker | 2024-01-01 |  | Yes | 100 | 80 |  | 0.5 | 5000
Patient Entrance Door | 0.05 | worker