## Without Timer
TEST: ACCURACY OF OPERATING POTENTIAL
Tolerance Sign | ±
Tolerance Value (kVp) | 5
Applied kVp | mA 1 | mA 2
60 | 60.1 | 60.2
80 | 80.1 | 80.2
100 | 100.1 | 100.2
120 | 120.1 | 120.2
TEST: TOTAL FILTRATION
Total Filtration Measured (mm Al) | 2.1
Total Filtration Required (mm Al) | 2
Total Filtration At kVp | 80
TEST: LINEARITY OF mAs LOADING
FDD (cm) | kV | mAs Range | Measured mR 1 | Measured mR 2 | Measured mR 3
100 | 80 | 5 | 4.1 | 4.2 | 4.1
100 | 80 | 10 | 8.1 | 8.2 | 8.1
100 | 80 | 20 | 16.1 | 16.2 | 16.1
100 | 80 | 50 | 40.1 | 40.2 | 40.1
Tolerance Operator | <=
Tolerance Value (CoL) | 0.1
TEST: CONSISTENCY OF RADIATION OUTPUT
Tolerance Operator | <=
Tolerance Value (CoV) | 0.05
FDD (cm) | Test kV | Test mAs | Meas 1 | Meas 2 | Meas 3
40 | 120 | 100 | 50.1 | 50.2 | 50.1
TEST: RADIATION LEAKAGE LEVEL
Distance | kV | mA | Time | Workload | Tolerance Value | Tolerance Operator | Tolerance Time | Location | Front | Back | Left | Right | Top
100 | 120 | 100 | 1 | 500 | 1 | <= | 1 | Tube | 0.01 | 0.02 | 0.01 | 0.01 | 0.02
TEST: RADIATION PROTECTION SURVEY REPORT
kV | mA | Time | Workload | Location | mR/hr
80 | 100 | 0.5 | 5000 | Control Console (Operator Position) | 0.02
80 | 100 | 0.5 | 5000 | Outside Patient Entrance Door | 0.01