## Without Timer
TEST: ACCURACY OF OPERATING POTENTIAL
Tolerance Sign | ±
Tolerance Value (kVp) | 5
Applied kVp | mA 1 | mA 2 | Average kVp | Remarks
60 | 60.2 | 60.5
70 | 70.1 | 70.4
80 | 80.3 | 80.1
TEST: TOTAL FILTRATION
Parameter | Measured (mm Al) | Required (mm Al) | At kVp
TotalFiltration | 2.6 | 2.5 | 80
TEST: LINEARITY OF mAs LOADING
FDD (cm) | 100 | kV | 80
mAs Range | Measured Output 1 | Measured Output 2 | Measured Output 3
5 | 0.5 | 0.51 | 0.49
10 | 1 | 1.01 | 0.99
20 | 2.5 | 2.51 | 2.49
Tolerance Operator | <=
Tolerance Value (CoL) | 0.1
TEST: CONSISTENCY OF RADIATION OUTPUT
FDD (cm) | 100
kVp | mAs | Output 1 | Output 2
80 | 10 | 0.5 | 0.51
Tolerance Operator | <=
Tolerance Value (CoV) | 0.05
TEST: RADIATION LEAKAGE LEVEL FROM X-RAY TUBE HOUSE
kV | 120 | mA | 10 | Time | 1.0 | Workload | 100 | Tolerance | 1.0 | Tolerance Operator | <
Location | Left | Right | Top | Up | Down | Max Leakage | Unit | Remark
Tube | 0.01 | 0.015 | 0.012 | 0.011 | 0.010 |  | mGy/h
Collimator | 0.012 | 0.011 | 0.014 | 0.013 | 0.012 |  | mGy/h
TEST: RADIATION PROTECTION SURVEY REPORT
Applied Current (mA) | 10 | Applied Voltage (kV) | 80 | Exposure Time (s) | 1.0 | Workload (mA.min/week) | 100
LOCATION | MAX. RADIATION LEVEL (MR/HR) | MR/WEEK | STATUS | RESULT
Control Console | 0.05 |  |  | Worker
Waiting Area | 0.02 |  |  | Public