## Sheet1
TEST: ACCURACY OF OPERATING POTENTIAL (KVP)
FFD (cm) | Tolerance Sign | Tolerance Value | Header 1 | Header 2 | Header 3 | Applied kVp | 22 | 33 | 11
100 | > | 2 | 22 | 11 | 12 | 60 | 60.2 | 60.1 | 60.3
[6 empty] | 80 | 80.1 | 80.2 | 80
[6 empty] | 100 | 100.2 | 100.1 | 100.3
TEST: TOTAL FILTRATION
Measured (mm Al) | Required (mm Al) | At kVp
2.5 | 2.5 | 80
TEST: OUTPUT CONSISTENCY
FFD (cm) | Tolerance Operator |  | Tol Value | Header 1 | Header 2 | Header 3 | kV | mAs | 34 | 22 | 11
100 | <= |  | 0.05 | 34 | 12 | 22 | 60 | 10 | 1.2 | 1.3 | 1.25
[7 empty] | 80 | 20 | 2.5 | 2.6 | 2.55
TEST: CENTRAL BEAM ALIGNMENT
FCD (cm) | kV | mAs | Observed Tilt |  | Tol Value | Tol Operator
100 | 80 | 20 | 0.5 |  | 1.5 | <=
TEST: CONGRUENCE OF RADIATION
Dimension | Observed Shift | Edge Shift | Percent FED | Tolerance
100 | 0.5 | 1 | 0.5 | 2
200 | 1 | 2 | 0.5 | 2
TEST: EFFECTIVE FOCAL SPOT
FCD (cm) | Focus Type | Stated Nominal | Measured Nominal
100 | Small | 0.6 | 0.62
| Large | 1.2 | 1.22
TEST: LINEARITY OF MAS LOADING
FDD (cm) | kV | Tol Operator | Tol Value | Header 1 | Header 2 | Header 3 | mAs | 90 | 100 | 123
100 | 80 | <= | 0.1 | 90 | 100 | 12 | 10 | 1.2 | 1.3 | 1.25
[7 empty] | 20 | 2.5 | 2.6 | 2.55
TEST: TUBE HOUSING LEAKAGE
FDD (cm) | kV | mA | Time (sec) | Workload | Tol Value | Tol Operator | Location | Left | Right | Front | Back | Top
100 | 80 | 100 | 1 | 5000 | 1 | <= | Tube | 0.5 | 0.3 | 0.4 | 0.2 | 0.6
[7 empty] | Collimator | 0.3 | 0.4 | 0.5 | 0.3 | 0.7
TEST: RADIATION PROTECTION SURVEY
Survey Date | Has Valid Calibration | Applied Current (mA) | Applied Voltage (kV) | Exposure Time (s) | Workload (mA min/week)
| Yes | 100 | 80 | 1 | 5000
LOCATION | MAX. RADIATION LEVEL (MR/HR) | CATEGORY
Control Room | 0.5 | worker
Corridor | 0.2 | public
TEST: HIGH CONTRAST SENSITIVITY
Measured (lp/mm) | Recommended Standard
5 | 4
TEST: LOW CONTRAST SENSITIVITY
Smallest Hole Size (mm) | Recommended Standard
2 | 1.5
TEST: ALIGNMENT TEST
Test Name | Sign | Value
Positional/re-positional accuracy of imager | <= | 1 mm
Alignment accuracy of X-ray beam | >= | 2 mm