## Without Timer
TEST: ACCURACY OF OPERATING POTENTIAL
Tolerance Sign | ±
Tolerance Value (kVp) | 5
Applied kVp | mA 1 | mA 2
60 | 60.2 | 60.5
70 | 70.1 | 70.4
80 | 80.3 | 80.1
TEST: TOTAL FILTRATION
Parameter | Measured mm Al | Required mm Al | At kVp
TotalFiltration | 2.6 | 2.5 | 80
TEST: LINEARITY OF mAs LOADING
FDD (cm) | 100 | kV | 80
Tolerance Operator | <=
Tolerance Value (CoL) | 0.1
mAs Range | Measured Output 1 | Measured Output 2 | Measured Output 3
12 | 1.2 | 1.25 | 1.22
50 | 2.4 | 2.45 | 2.42
TEST: CONSISTENCY OF RADIATION OUTPUT
Tolerance Operator | <=
Tolerance Value (CoV) | 0.05
FDD (cm) | 100
kVp | mAs | Output 1 | Output 2 | Output 3
80 | 100 | 10.2 | 10.1 | 10.3
TEST: RADIATION LEAKAGE LEVEL FROM X-RAY TUBE HOUSE
FDD (cm) | 100 | kV | 120 | mA | 10 | Time | 1 | Tolerance | 1
Workload | 100
Location | Front | Back | Left | Right | Max | Unit | Remark
Tube | 0.01 | 0.012 | 0.015 | 0.011 |  | mGy/h
Collimator | 0.012 | 0.014 | 0.011 | 0.013 |  | mGy/h
TEST: RADIATION PROTECTION SURVEY REPORT
|  | Applied Current (mA) | 10 | Applied Voltage (kV) | 80 | Exposure Time (s) | 1 | Workload (mA.min/week) | 100
LOCATION | MAX. RADIATION LEVEL (MR/HR) | RESULT
Control Console | 0.05 | Worker
Waiting Area | 0.02 | Public