## Without Timer
TEST: ACCURACY OF OPERATING POTENTIAL (kVp)
Tolerance Sign | -
Tolerance Value (kVp) | 2
Total Filtration Required (mm Al) | 2.5
Total Filtration At kVp | 80
Applied kVp | 50 mA | 100 mA
60 | 59.8 | 60.2
80 | 79.5 | 80.5
100 | 99.2 | 100.8
120 | 119.5 | 120.5
TEST: OUTPUT CONSISTENCY
FDD (cm) | 100
kVp | mAs | Meas 1 | Meas 2 | Meas 3 | Meas 4 | Meas 5
80 | 10 | 0.85 | 0.84 | 0.86 | 0.85 | 0.85
Tolerance Operator | <=
Tolerance (CoV) | 0.05
TEST: CENTRAL BEAM ALIGNMENT
FFD (cm) | 100 | kV | 80 | mAs | 10
Observed Tilt (cm) | 0.5
Tolerance (cm) | 1.5
TEST: CONGRUENCE OF RADIATION
FFD (cm) | 100 | kV | 80 | mAs | 10
Dimension | Observed Shift (cm) | Edge Shift (cm) | Tolerance (%)
Length | 0.2 | 0.1 | 2
Width | 0.3 | 0.1 | 2
TEST: EFFECTIVE FOCAL SPOT
FFD (cm) | 60
Focus Type | Stated Focal Spot of Tube (f) | Measured Focal Spot (Nominal)
Small | 0.6 | 0.65
Large | 1.2 | 1.3
TEST: LINEARITY OF mAs LOADING STATIONS
FDD (cm) | 100 | kV | 80
mAs Applied | Meas 1 | Meas 2 | Meas 3
5 | 0.43 | 0.42 | 0.43
10 | 0.86 | 0.85 | 0.87
20 | 1.72 | 1.71 | 1.73
Tolerance Operator | <=
Tolerance (CoL) | 0.1
TEST: RADIATION LEAKAGE LEVEL
FDD (cm) | 100 | kV | 125 | mA | 20 | Time (min) | 2 | Workload | 20
Tolerance Value (mGy/h) | 1
Tolerance Operator | <
Location | Left | Right | Front | Back | Top
Tube | 0.05 | 0.04 | 0.06 | 0.05 | 0.04
Collimator | 0.02 | 0.03 | 0.02 | 0.03 | 0.02